## Slide 1: Quality Control of GEM-Foils with Optical Scanning System
Matti Kalliokoski
Timo Hildén
3rd JointGEM Meeting
Helsinki 15/7/2010

## Slide 2: XYZ-linears
Movement of the linears 1 m in X and Y direction
  Maximum velocity ~5.5 cm/s
  Accuracy < 100 μm over 1 m
  Repeatability < 10 μm
Z movement 12 cm
  Space between z-linear and top of the table ~4.5 cm
  Maximum velocity 4 cm/s

## Slide 3: Camera
Artray Artcam-900MI – USB 2.0
9 Mpix color camera with c-mount
Maximum resolution of 3488 x 2616 pixels
  Pixel size of 1.75 x 1.75 μm
With maximum resolution frame rates from 1.2 to 3.9 per second

## Slide 4: Optics
Two telecentric lenses with magnifications of 1.00 and 0.37
Image sizes respectively 6.1 x 4.6 and 16.5 x 12.4 mm
Pixel sizes 1.75 and 4.73 μm
Working distance ~63 mm

## Slide 5: Laser Telemeter
Focal distance is adjusted automatically during the scans
Distance is measured with Micro-Epsilon OptoNCDT-1302 -20 laser telemeter
Maximum static resolution 4 μm, dynamic 10 μm
Measuring rate 750 Hz
Tilted in angle to measure from the focal point of the camera

## Slide 6: Lighting
Imagon 95x95 cm LED light table for background light
  Unadjustable
  Brightness of the holes can be adjusted only by exposure time
  In some cases diffusors can be used
    Not for GEMs
HXD-IA Dimming Ring LED Illuminator attached to the camera for foreground light
  Adjustable
  Different foils require different intensities

## Slide 7: Vibration Control
Low frequency tremble was observed to disturb the scanning
To reduce problems on keeping the focus during the scans, two piezo elements were fixed to the end of the X-linear to measure the vibrations
From measurements we could get the optimal waiting time between the movement and image taking

## Slide 8: Scanning of GEM-Foils
LabView program controls the movement of the components
With maximum resolution scan of an area of 30 x 10 cm takes about 4 hours
  Scan of the whole area of 95 x 95 cm takes about 30 hours
From an area of 30 x 10 cm about 1500 images are taken
  Average size of a sharp image is about 20 MB
  => 1500 images = 30 GB
  => Online analysis with data reduction is required

## Slide 9: Distorsions
Framed foils generally have some tensions on the surface
This will distort the frames
Straightening the foil ususally creates local undulations
This can be seen also with unframed foils but with smaller distorsions

## Slide 10
Foil Types

## Slide 11: Foil Types

## Slide 12: Foil Spectrum

## Slide 13: Foil Spectrum

## Slide 14: Mirror Foils

## Slide 15: Mirror Foils

## Slide 16: Analysis

## Slide 17: Defects
Solvent left on the surface
Incompletely etched holes or done only from one side
Etching defects, Kapton or copper fully removed

## Slide 18: Other Studies

## Slide 19: Conclusions
We have obtained an optical scanning system for quality control and studies of GEM-foils
The setup can scan foils with maximum resolution of 1.7 micrometers per pixel
Different defects can be discovered
Analysis tools have been developed to study the quality of the foils
With preliminary analysis mean inner hole size of 49 ± 6 μm has been measured from TOTEM foils